# Where an unaided clinician’s deep-learning study breaks: a three-rung external-validation ladder for spleen segmentation, and a preprocessing contract that fails in silence

Demonstration Author

2026-07-31

## Abstract

Background and objective. Agentic toolkits can now scaffold a medical-imaging model, and a clinician-researcher without an engineer can plausibly reach a trained network. Where that path breaks is not known.

Methods. We trained nnU-Net v2 (1) for 3-D spleen segmentation on 32 cases of the Medical Segmentation Decathlon Task09 (2) under a patient-level split, and evaluated it on three labelled rungs: the internal held-out set, AMOS22 CT (n = 300) as a genuinely external cohort, and AMOS22 MRI (n = 60) as a modality shift (3). Metrics, the handling of cases containing no spleen, the subgroup cut-points and a written prediction for the MRI rung were fixed before any prediction existed; Dice is reported with a boundary metric and never accuracy (4).

Results. Internal median Dice was 0.9595 (95% CI 0.9367 to 0.9734) and fell to 0.8932 (0.8633 to 0.9108) on external CT, a difference of −0.0662 (−0.0996 to −0.0416). Performance was lowest where the task matters most: enlarged spleens scored 0.7694 against 0.9114 for normal volumes. On MRI, median Dice was 0.0152 (0.0000 to 0.0626), and the pipeline reported no error while returning a file for all 60 cases, 20 of them empty. Replaying the trained plan’s normalisation located an incompatibility present in every MRI case and no CT case: all 300 CT cases contain negative voxels and 2.7% of voxels exceed the clip ceiling, while no MRI case contains a negative voxel and a median 23.2% exceed it. A counterfactual arm that changed only the input intensity scale — same checkpoint, same folds, no retraining — recovered median Dice to 0.3016 (0.1744 to 0.4048), a difference of +0.2864 (+0.1204 to +0.4048), while remaining −0.5916 (−0.7259 to −0.4674) below the external CT arm.

Conclusion. Both mechanisms are real and their sizes differ: of the collapse observed against the internal arm, roughly 0.29 Dice is attributable to the preprocessing contract and roughly 0.59 to a representation that does not transfer. The third rung is a constructed test — the plan named the normalisation contract and predicted the collapse before inference ran — so it demonstrates that the pipeline stays silent about a known incompatibility, not that the failure was discovered in unaided work. Within that limit it shows a defect class worth tooling against — one that produces a number rather than an error, and that the toolkit’s own profiler had already flagged at Minor severity in a directory no later step reads. Generalisation beyond this example is a hypothesis, not a result.

## Introduction

The practical question behind agentic research tooling is not whether a clinician can obtain a trained model. They can. It is whether the study that results is defensible, and where it stops being so. Answering that requires running the whole path, including the tedious parts, and recording the failures rather than the finished artifact.

What kind of evidence this is. Rungs 1 and 2 are ordinary internal and external validation and are reported as such. Rung 3 is different and should be read differently: while writing the evaluation plan we inspected the trained configuration, found that it carries a CT intensity transform, and predicted in writing that MRI performance would collapse for that reason. We then ran the arm anyway. That makes rung 3 a constructed demonstration — a stress test of whether the pipeline says anything when handed an input it cannot process — and not a discovery made in the course of unaided work. Read as a discovery it would overclaim; read as a constructed test it answers a narrow question cleanly, which is: given a known incompatibility, does anything in the workflow surface it? The answer is no, and the run exits 0.

The failure modes this literature documents most often are about what the model learned: a pneumonia classifier that encoded the scanner rather than the disease (5), a hip-fracture model driven by healthcare process variables (6). The methodological guidance that follows from them concerns cohort design and external testing (7). The failure recorded here is a different species. It is not about what the model learned; it is about what the pipeline silently did to the images before the model ever saw them.

We chose 3-D spleen segmentation because it admits a clean ladder of increasing distributional distance with ground truth at every step, and because both datasets download without registration. Access, not narrative quality, decided the choice. ACDC to M&Ms is the better domain-shift pair and sits behind registration walls; a demonstration about reproducibility cannot open with a step the reader is unable to take.

Two properties of the setup have to be stated before any result is read. First, nnU-Net won the 2018 Medical Segmentation Decathlon and its self-configuring heuristics were developed against that benchmark family (1,2), so the internal rung on MSD Task09 demonstrates that the pipeline runs and does not independently evidence the method. AMOS22 postdates nnU-Net, so the external rungs carry the evidence. Second, a clinician sourcing a model from a public repository checks its licence and its citation count, and neither of those surfaces the first point. That benchmark-provenance hazard is a known one in this field: competition rankings are sensitive to design choices that a leaderboard position does not expose (8), and nnU-Net’s own authors have since argued for more rigorous validation practice around it (9).

## Methods

Data. MSD Task09 Spleen (CC-BY-SA 4.0) (2), 41 labelled cases of 61. AMOS22 (CC BY 4.0) (3), 300 labelled CT and 60 labelled MRI of 600, downloaded as the Zenodo data record doi:10.5281/zenodo.7262581, which is what carries the licence. Dataset facts were read from the shipped files rather than from the accompanying prose: the spleen index is 1 in both label sets, the CT label set runs 0 to 15 and the MRI set 0 to 13, and the readme’s modality rule is off by one against a contiguous identifier space. Cohort characteristics by arm are in Table 1.

Split and preprocessing. A patient-level split with seed 42 assigned 32 cases to training and 9 to the held-out set. Disjointness was proved by set arithmetic rather than asserted. nnU-Net fits its resampling target and its intensity statistics from whatever occupies its raw-data directory, so the fit scope is decided by which files are copied there and not by any flag. Held-out cases were therefore physically excluded from that directory, and the dataset builder aborts if one appears. The counterfactual, in which every case is present and the split is applied afterwards, is retained as a declared manifest and fires three major preprocessing-leakage findings under the same gate that passes the real one.

Training and inference. nnU-Net v2 (1), configuration 3d_fullres, five folds, 100 epochs per fold. Measured epoch time of 360 to 417 s makes nnU-Net’s default 1000 epochs approximately 21 days for five folds on the available hardware, so the reduced schedule follows from arithmetic and is disclosed wherever a number from this model appears. The epoch time, the derived 21-day figure and the approximately 50 GPU-hours were read from the training job logs on the cluster; those logs are not shipped with this demo, so unlike every other number here they cannot be re-derived from the repository. Five folds were retained so that the deviation from default remains a single change. Inference used the five-fold ensemble at nnU-Net’s default settings with mirroring test-time augmentation enabled, on all three arms.

Evaluation. Metrics were fixed before any prediction existed and follow Metrics Reloaded (4): Dice is reported with HD95, since a boundary metric is required alongside an overlap metric, and accuracy is never reported because the target occupies 0.2 to 0.4% of the volume. HD95 was computed with the supplied voxel spacing, which matters here because the volumes are anisotropic.

HD95 has a smaller denominator than Dice, and the difference is not random. A non-empty reference with an empty prediction has a Dice of 0 but no predicted surface, so no symmetric surface distance exists for it. The case is left blank rather than imputed, because the usual imputation — the image diagonal — is a made-up number. Those are the cases that failed hardest, so an HD95 median quoted against the Dice denominator is optimistic by exactly the worst cases, and the gap widens with the failure rate: HD95 exists for 9 of 9 internal cases, 270 of 298 scored external CT cases, and 40 of 59 scored MRI cases. Every HD95 below is therefore reported with its own n (Table 2), and the MRI value in particular is a statistic over the 40 cases that produced a surface at all. HD95 here is therefore a success-conditioned estimand: it describes boundary agreement among the cases that produced a boundary, and it should be read together with the empty-prediction count for the same arm rather than as that arm’s boundary performance.

The primary summary is the median, and that is a choice with consequences. On the external CT arm the mean Dice is 0.7116 against a median of 0.8932, so which one leads changes the apparent result. The estimand here is typical-case performance — what the model does on a case drawn from the cohort — for which the median is the appropriate summary and is resistant to the 28 empty predictions. The mean is the expected per-case performance and absorbs those failures; it is the better summary for anyone budgeting total error. Both are reported side by side in Table 2 for exactly that reason, and the empty-prediction count is reported with them so the median is never read as if the failures had not happened.

Uncertainty. Every median is reported with a percentile bootstrap 95% interval over cases — 10,000 resamples, seed 20260725, seeded per (arm, metric) so that an arm’s interval does not depend on how many other arms were bootstrapped before it. Between-arm differences carry their own interval: a difference of two separately estimated medians is not an inferential quantity on its own, and the internal arm’s n = 9 contributes most of the uncertainty, so each Δ-Dice is a bootstrap interval for the difference in population medians with both arms resampled independently at the case level.

Three of the 360 labelled AMOS cases contain no spleen voxel. Dice is undefined on an empty reference, so those cases were excluded from the metric distributions, which fixes the denominators at 298 of 300 and 59 of 60, and were reported separately with their predicted volume so that a hallucinated organ appears on its own line. Subgroup cut-points, at 100 and 250 mL of spleen volume and 2 and 5 mm of slice thickness, were fixed in advance and are identical across arms; a per-arm quantile would have made the arms non-comparable. Reporting follows CLAIM 2024 (10,11), with TRIPOD+AI (12) consulted for the prediction-model items.

Provenance of the plan, stated precisely. EVALUATION_PLAN.md was written before inference was run and is shipped unedited, but this repository cannot prove that ordering: the plan and the rung-3 per-case results first appear in the same commit, because the demo was committed in one piece. A reader should treat the chronology as documented by the authors rather than as independently verified. What the repository does carry independently is the profiler’s own output (qc/amos22_dataset_profile_spleen.json), which contains the INTENSITY_SCALE_INCONSISTENT claim and its Minor severity as machine-readable fields.

The counterfactual arm (rung 3b). Measuring what the normaliser does to an image does not show that it caused the failure. Separating the preprocessing contract from a representation that does not transfer needs one more arm, in which exactly one thing changes: the intensity scale of the input. Each MRI volume was affinely mapped so that the 0.5th-to-99.5th percentile range of its in-body voxels (> 0; AMOS MRI air is exactly 0) lands on the CT fingerprint’s foreground window, read from the trained plans.json. The normaliser is not bypassed or replaced — it is handed input in the domain it assumes, so that on MRI it performs the operation it performs on CT. After the mapping, the share of voxels clipped at the ceiling falls from a median of 23.2% to 0.4%. The checkpoint, the five folds, the ensemble, the test-time augmentation, the ground truth and every metric rule are unchanged, and nothing is retrained. The arm, its rationale, the rejected alternatives and a written prediction were fixed in COUNTERFACTUAL_PLAN.md before it was run.

Normalisation evidence. To characterise what the trained plan does to each arm, the normalisation recorded in the trained plans.json was replayed over every image of both external arms, measuring per case the presence of negative voxels, the fraction of voxels above the clip ceiling, and the number of intensity levels surviving the clip.

## Results

The ladder. Internal median Dice was 0.9595 (95% CI 0.9367 to 0.9734), with HD95 1.78 mm over all 9 cases. On external CT, 298 of 300 cases were scored, with median Dice 0.8932 (0.8633 to 0.9108) and HD95 5.68 mm over the 270 cases that produced a predicted surface; the difference in population medians against the internal arm is −0.0662 (95% CI −0.0996 to −0.0416). On MRI, 59 of 60 cases were scored, with median Dice 0.0152 (0.0000 to 0.0626), HD95 70.05 mm over 40 of those 59, and a difference of −0.9443 (−0.9715 to −0.8813). The counterfactual arm (rung 3b) scored 0.3016 (0.1744 to 0.4048) on the same 59 cases (Table 2). Figure 1 shows the per-case distributions behind those medians and Figure 2 the across-cohort comparison with its intervals. For context, nnU-Net’s own cross-validation over the 32 training cases gave fold Dice values of 0.8437, 0.9304, 0.9659, 0.9696 and 0.9507.

What the headline hides. Twenty-eight of the 300 external CT predictions were empty, as were 20 of the 60 MRI predictions. Both spleen-free CT cases received a non-empty prediction. Figure 1 makes the reason plain: the external CT arm carries a long tail to zero, and the MRI arm has most of its mass at the floor with a handful of cases reaching 0.91. In the pre-specified subgroups (Table 3, Figure 3), external CT performance was lowest on enlarged spleens above 250 mL, at 0.7694 (n = 67), against 0.9114 for the normal 100 to 250 mL range (n = 194) and 0.8157 for volumes below 100 mL (n = 37). Thick slices of 5 mm or more scored 0.8752 (n = 219) against 0.9292 for the 2 to 5 mm range (n = 74). The thin-slice stratum has n = 5 and the thick MRI stratum is empty; both are descriptive and cannot support a stable contrast. These are stratum estimates with their own intervals, not a tested interaction; no interaction test is reported and none is implied.

A preprocessing incompatibility is present in every MRI case, and no CT case. Two rival explanations were tested and rejected before the result was accepted. The label index is correct: spleen is 1 in the MRI label set, and the resulting reference volumes are physiological, with a median of 187.6 mL across the 59 scored cases (Table 1; 186.8 mL if the target-free case is included). The images are genuinely MRI, with a minimum of 0 and no air floor.

The trained plan applies CTNormalization with the training fingerprint’s parameters, clipping to [-38, 174] HU and z-scoring by mean 89.8 and SD 39.8. Replaying it over both arms separates them cleanly (Table 4, Figure 4). All 300 CT cases contain negative voxels, a median of 2.7% of voxels exceed the clip ceiling (range 1.0 to 6.9%), and 213 intensity levels survive. No MRI case contains a negative voxel, a median of 23.2% of voxels exceed the ceiling (range 0.3 to 68.6%), and the surviving levels fall to a median of 175 with a minimum of 2. Hounsfield units are defined by an air floor near -1000, so a volume that never goes negative is not in that unit system, and the clip window is not a soft-tissue window there. It is an arbitrary slice near the bottom of an arbitrary-unit range. Case accounting for both source datasets, including every exclusion and the three target-free cases, is in Figure 5.

The counterfactual separates the two mechanisms, and both are real. Rescaling the MRI into the plan’s intensity domain — one change, no retraining — raised median Dice from 0.0152 to 0.3016 (95% CI 0.1744 to 0.4048), a difference in population medians of +0.2864 (+0.1204 to +0.4048), and reduced empty predictions from 20 to 15 of 60. The same arm remains −0.5916 (−0.7259 to −0.4674) below external CT. Neither interval crosses zero.

So of the −0.9443 observed against the internal arm, roughly 0.29 Dice is attributable to the preprocessing contract and roughly 0.59 to a representation that does not transfer to this modality. A twenty-fold recovery from an input rescaling alone is not a rounding effect: for much of the original collapse, the network was not failing to read MRI so much as never receiving it. Neither is the residual: a correctly-scaled MRI still loses most of the performance the CT cohort retains, because an affine map restores dynamic range and cannot make an MR sequence’s tissue contrast agree with the ordering a Hounsfield window encodes.

This is a decomposition, not the identification of a single cause, and it is bounded by its own design: it compares two input scalings of one model on one modality pair.

No part of the run signals this. The inference command has no argument that declares the modality of the incoming images, and the normaliser travels with the checkpoint. The dataset’s own dataset.json declares a single CT modality for a collection that contains 100 MRI volumes, so a pipeline that consults that field reaches the same wrong answer.

## Discussion

The clinician in this study did the things the methodological literature asks for (7). The licence was checked, the model’s provenance conflict with the internal benchmark was identified and disclosed, the split was patient-level and seed-locked, disjointness was proved rather than assumed, the external cohort was genuinely external, the metrics were appropriate to the task (4), the cases without a target were handled by a rule fixed in advance, and the subgroups were pre-specified. The third rung still produced a meaningless number, and a reader without ground truth would have had nothing to tell them so.

That last clause is the whole claim, and it is narrower than it first looks. We are not claiming to have discovered this failure by walking the unaided path: the evaluation plan named the normalisation contract and predicted the collapse before inference ran (see the Introduction). What the run demonstrates is the silence — that a pipeline handed an input it cannot process completes, exits 0, and reports a number, with nothing in the command, the logs or the output signalling the problem. That is a defect class worth tooling against, and it is one this study constructed rather than encountered.

The toolkit’s own record complicates that account, and in a way worth stating precisely. Before any training, the dataset profiler returned INTENSITY_SCALE_INCONSISTENT on AMOS22: “500/600 cases bottom out near air (<= -500) and the rest do not, mixed modality, or a rescale not applied to part of the cohort”. That is a true positive, it names the exact property that later broke the MRI rung, and it is in the shipped quality-control output, where the claim and its severity are machine-readable fields. It was filed as Minor.

So detection was not the missing piece. Routing and severity were. Nothing carries a profiling-stage observation forward to the inference stage, nothing compares the intensity domain a trained plan assumes against the arm it is about to be applied to, and a finding whose cost turns out to be a Dice of 0.015 was ranked below the threshold at which anyone stops. A gate that fires correctly into a directory no later step reads is, operationally, a gate that did not fire. We offer that as a hypothesis this example motivates, not as a property established by it — one organ, one architecture, one dataset pair and one deliberately chosen mismatch cannot support a general claim about tooling, and no independent user ever walked this path.

The written prediction called the mechanism, then overstated it. Before any prediction existed, the evaluation plan named the normalisation contract as the thing that would break the MRI rung, and the outcome matched. (Its chronology is documented by the authors rather than provable from this repository; see Methods.) It also predicted that the clipped volume would retain “only two values”, reasoning from a single MRI case with an unusually large intensity range. Measured across the cohort, two levels is the extreme rather than the norm: the median is 175 levels and 23.2% of voxels clipped. Generalising a mechanism from one case overstated it, and the plan is left uncorrected in the repository with this note beside it.

Limitations. The training cohort is 32 cases from a single institution and only one nnU-Net configuration was trained, so this is nnU-Net the network rather than nnU-Net the method (9). Subgroup estimates are external-only because 9 internal cases cannot support them, and whether the spleen-free cases reflect splenectomy or annotation omission was not determined. The slice-thickness axis is the third voxel dimension, which is the through-plane axis for every CT case but not for 26 of the 60 MRI volumes; nine scored MRI cases would change stratum under a coarsest-axis rule, so the MRI subgroup rows should be read with that in mind.

The counterfactual arm separates preprocessing failure from representation failure but does not exhaust the question: percentile matching is the weakest intervention that fixes the declared defect, and a volume-wise z-score — what nnU-Net would have chosen had the dataset been declared MR — is a different counterfactual that was not run. No model trained for multi-modality use was attempted here (13). Finally, this is one worked example authored by the people who built the toolkit — no independent users, no comparator workflow, no repetition.

## Conclusion

A clinician-researcher without an engineer reached a defensible external-validation result on CT. On a third arm, constructed to carry a known incompatibility, the same pipeline returned output for every case, exited 0, and reported a number that means nothing — and the toolkit’s own profiler had already recorded the underlying property, at Minor severity, in a directory no later step reads.

The change that example suggests is not another detector but a route: carry a profiling-stage observation to the point of use, and compare the intensity domain a trained plan assumes against the arm it is about to be applied to, before an inference run is allowed to report a number. Whether that would prevent failures beyond the one constructed here is untested, and this study cannot settle it — one organ, one architecture, one dataset pair, one author team. What it does establish is that the failure can be complete, and completely quiet, at the same time.

## Tables

Table 1. Cohort characteristics by evaluation arm: case counts, target-free counts, spleen volume and slice thickness.

Table 2. Performance by arm, with the target-free, empty-prediction and HD95 denominators reported beside each estimate.

Table 3. Pre-specified subgroups, identical cut-points across arms.

Table 4. What the trained plan’s normaliser does to each external arm.

## Figure legends

Figure 1. Per-case Dice by cohort. Box plots with jittered per-case points; the spread is what the medians in Table 2 conceal.

Figure 2. Median Dice with bootstrap 95% confidence intervals across the three cohorts, with the internal arm as the dashed reference.

Figure 3. Pre-specified subgroups on the external CT arm, both axes with identical cut-points.

Figure 4. Why rung 3 collapsed, measured on both external arms. (a) Per-case minimum intensity: every CT case has a Hounsfield air floor near -1000; no MRI case has one. (b) The share of each volume flattened at the trained plan’s clip ceiling under the same normaliser.

Figure 5. Case accounting across both source datasets and the three rungs, including the three cases excluded from the metric distributions because their reference contains no spleen.

## Data and code availability

Both datasets are public and are not vendored here: MSD Task09 Spleen (CC-BY-SA 4.0) and AMOS22 (CC BY 4.0, Zenodo 7262581). The pipeline, the per-case results behind every number, the deterministic gate outputs including the leakage counterfactual, and the friction log are in demo/05_msd_amos_spleen/. reproduce.sh re-runs the gates and the full across-cohort analysis on a laptop; training requires approximately 50 GPU-hours and is documented rather than executed.

## References

1. 	Isensee F, Jaeger PF, Kohl SAA, Petersen J, Maier-Hein KH. nnU-net: A self-configuring method for deep learning-based biomedical image segmentation. Nature Methods. 2021;18:203–211. doi: 10.1038/s41592-020-01008-z.

2. 	Antonelli M, Reinke A, Bakas S, et al. The medical segmentation decathlon. Nature Communications. 2022;13. doi: 10.1038/s41467-022-30695-9.

3. 	Ji Y, Bai H, Ge C, Yang J, et al. AMOS: A large-scale abdominal multi-organ benchmark for versatile medical image segmentation. Advances in neural information processing systems 35 (NeurIPS 2022). 2022. p. 36722–36732. doi: 10.52202/068431-2661.

4. 	Maier-Hein L, Reinke A, Godau P, et al. Metrics reloaded: Recommendations for image analysis validation. Nature Methods. 2024;21:195–212. doi: 10.1038/s41592-023-02151-z.

5. 	Zech JR, Badgeley MA, Liu M, et al. Variable generalization performance of a deep learning model to detect pneumonia in chest radiographs: A cross-sectional study. PLOS Medicine. 2018;15:e1002683. doi: 10.1371/journal.pmed.1002683.

6. 	Badgeley MA, Zech JR, Oakden-Rayner L, et al. Deep learning predicts hip fracture using confounding patient and healthcare variables. npj Digital Medicine. 2019;2. doi: 10.1038/s41746-019-0105-1.

7. 	Park SH, Han K. Methodologic guide for evaluating clinical performance and effect of artificial intelligence technology for medical diagnosis and prediction. Radiology. 2018;286:800–809. doi: 10.1148/radiol.2017171920.

8. 	Maier-Hein L, Eisenmann M, Reinke A, et al. Why rankings of biomedical image analysis competitions should be interpreted with care. Nature Communications. 2018;9. doi: 10.1038/s41467-018-07619-7.

9. 	Isensee F, Wald T, Ulrich C, et al. nnU-net revisited: A call for rigorous validation in 3D medical image segmentation. Medical image computing and computer assisted intervention (MICCAI 2024). 2024. p. 488–498. doi: 10.1007/978-3-031-72114-4_47.

10. 	Tejani AS, Klontzas ME, Gatti AA, et al. Checklist for artificial intelligence in medical imaging (CLAIM): 2024 update. Radiology: Artificial Intelligence. 2024;6(4). doi: 10.1148/ryai.240300.

11. 	Mongan J, Moy L, Kahn CE. Checklist for artificial intelligence in medical imaging (CLAIM): A guide for authors and reviewers. Radiology: Artificial Intelligence. 2020;2:e200029. doi: 10.1148/ryai.2020200029.

12. 	Collins GS, Moons KGM, Dhiman P, et al. TRIPOD+AI statement: Updated guidance for reporting clinical prediction models that use regression or machine learning methods. BMJ. 2024;e078378. doi: 10.1136/bmj-2023-078378.

13. 	Wasserthal J, Breit H-C, Meyer MT, et al. TotalSegmentator: Robust segmentation of 104 anatomic structures in CT images. Radiology: Artificial Intelligence. 2023;5. doi: 10.1148/ryai.230024.